INSTRUMENTAL METHODS OF ANALYSIS- BP701T

UNIT :2 Flame photometry

CLASS:4

TOPIC: FLAME PHOTOMETRY- PRINCIPLE, INTERFERENCES

INTRODUCTION

Flame photometry (Flame Atomic Emission Spectrometry) is a branch of spectroscopy in which the species examined in the spectrometer are in the form of atoms. A photoelectric flame photometer is an instrument used in inorganic chemical analysis to determine the concentration of certain metal ions among them sodium, potassium, calcium and lithium.  Flame Photometry is based on measurement of intensity of the light emitted when a metal is introduced into flame. The wavelength of colour tells what the element is (qualitative). The intensity of colour tells us how much of the element present (quantitative).

PRINCIPLE

The basic principle upon which Atomic Spectroscopy works is based on the fact that "Matter absorbs light at the same wavelength at which it emits light".  Atoms of elements   subjected to hot flame .So specific quantum of thermal energy  absorbed by orbital electrons becomes unstable at higher energy and release energy of photons by getting to ground state. When a metal solution is burned, it gives a coloured flame and each metal ion gives different coloured flame. Flame tests, therefore, can be used to test for the absence or presence of a metal ion.

Basic concept involved here is when Liquid sample containing metal salt solution is introduced into a flame,solvent is first vaporized, leaving particles of solid salt which is then vaporized into gaseous state. Gaseous molecule dissociate to give neutral atoms which can be excited (made unstable) by thermal energy of flame. The unstable excited atoms emit photons while returning to lower energy state. The measurement of emitted photons forms the basis of flame photometry. Under constant and controlled conditions, the light intensity of the characteristic wavelength produced by each of the atoms is directly proportional to the number of atoms that are emitting energy, which in turn is directly proportional to the concentration of the substance of interest in the sample.  Various metals emit a characteristic colour of light when heated.

INTERFERENCES

In determining the amount of a particular element present, other elements can also affect the result. Such interference may be of 4 kinds:-

- Spectral interferences: - There are several types of spectral interferences which are explained below.

- The first type of interference arises when two elements exhibit spectra, which partially overlap, and both emit radiation at some particular wavelength. The detector cannot distinguish between the sources of radiation and records the total signal, thus resulting in incorrect answer. Such interferences are more common at high flame temperatures because numerous spectral lines are produced at high temperatures. For example, the Fe line at 324.73 nm overlaps with the Cu line at 324.75 nm. Such interference can be overcome either by taking measurements at an alternative wavelength which has no overlap, if available, or by removing the interfering element by extraction. Alternatively, one may make a calibration curve, which is prepared from a solution having similar quantities of the interfering element. •

- The second type of spectral interference deals with spectral lines of two or more elements which are close but their spectra do not overlap. This type of interference becomes a problem when a filter is used as the device to isolate spectral lines. A filter may allow spectral lines separated by 5.0-10.0 nm to pass through, thus resulting in an error in the analysis. Such interferences can be reduced by increasing the resolution of the spectral isolation system. However, the interference cannot be eliminated entirely due to the finite width of the spectral isolation system and the finite slit width in such systems.

- A third type of spectral interference occurs due to the presence of continuous background which arises due to high concentration of salts in the sample, especially of alkali and alkaline earth metals. Some organic solvents also produce a continuous background. This type of interference can be corrected by using suitable scanning technique.

- Ionic interferences:

In some cases, high temperature flame may cause ionisation of some of the metal atoms, e.g., in case of sodium, it can be given as follows. Na -Na+ + e - The Na+ ion possesses an emission spectrum of its own with frequencies, which are different from those of atomic spectrum of the Na atom. This reduces the radiant power of atomic emission. This interference can be eliminated by adding a large quantity of a potassium salt to the standards as well as sample solutions. The addition of potassium salt suppresses the ionisation of sodium, as the potassium atom itself undergoes ionisation due to low ionisation energy. Thus, the sodium atom emission is enhanced.

This type of interference is restricted to alkali metals.

- Chemical interferences:- The chemical interferences arise out of the reaction between different interferents and the analyte which  Includes:

- Cation - anaion interference: – The presence of certain anions, such as oxalate, phosphate, sulfate, in a solution may affect the intensity of radiation emitted by an element. E.g., calcium and phosphate ion forms a stable substance, as Ca3(PO4)2 which does not decompose easily, resulting in the production of lesser atoms.

- Cation - cation interference: - In many cases, mutual interferences of cations have been observed, resulting in reduced signal intensity of the element being determined. These interferences are neither spectral nor ionic in nature Eg. aluminium interferes with calcium and magnesium

(d)Interference due to oxide formation

This type of interference arises due to the formation of stable metal oxide if oxygen is present in the flame, resulting in reduced signal intensity. The alkaline earth metals are subject to this type of interference. This type of interference can be eliminated by either using very high flame temperature to dissociate the oxides or by using oxygen-deficient environment to produce excited atoms.

- Other Factors

A number of other factors affect the intensity of light emission from a given sample solution. For example, the organic solvents in a sample may also influence the intensity of the emission line by changing the viscosity and surface tension of the liquid, which in turn alter the rate at which the sample is aspirated into the flame. They also affect the flame temperature through their contribution to the heat of combustion. An increase in the line intensity is usually observed.