INSTRUMENTAL METHODS OF ANALYSIS- BP701T

UNIT : 2 Atomic absorption spectroscopy

CLASS: 7

TOPIC: INSTRUMENTATION OF AAS

For all types of atomic absorption spectrometers, the following components are required

I. Source of radiation eg. Hollow cathode lamp

II. Chopper

III. Nebulizer-burner system (Atomization unit)

IV. Atomizer

V. Monochromators

VI. Detectors (Photometric measuring system)

VII. Amplifiers

VIII. Read out device  (Recorder)

Figure 7.1: Instrumentation of atomic absorption spectrophotometer

- Light source (Source of radiation )

The primary need of the light source is to generate a narrow line profile with a minimum background. It should also have a reliable and reproducible output with enough intensity to provide a high signal-to-noise ratio. In, atomic absorption spectroscopy generally, two types of light sources are used i.e. hollow-cathode lamps (HCL) and electrodeless discharge lamps (EDL).

Except for analytes such as arsenic and selenium, hollow-cathode lamps are the most typically employed, while electrodeless discharge lamps are used less frequently than HCLs.

- II. Chopper

The light source should be modulated and mechanically chopped to differentiate between the light from the source and the emission from the sample cell. Chopper distinguishes between absorption and emission signals. It is situated between the light source and the atomizer.

- III. Nebulizer – burner system

It produces a mist or aerosol of the test solution i.e. it converts the test solution to gaseous atoms It sucks in liquid samples at a controlled rate and produces a fine aerosol spray for use in the flame. It effectively combines aerosol, fuel, and oxidant.

- IV. Atomizer

Atomization is the process of separating particles into individual molecules and breaking molecules down into atoms. This is done by heating the analyte in a flame or graphite furnace to high temperatures.

To generate atoms from a sample, two systems are usually utilized. Aspiration involves sucking a sample solution into a flame, whereas electrothermal atomization involves placing a drop of sample into a graphite tube that is subsequently heated electrically. Some instruments include both atomization processes but only have one pair of lights. After selecting the suitable lamp, it is directed toward one or more atomization systems.

- V. Monochromator (Prisms Gratings)

A line isolation device is necessary to ensure that only light of a wavelength specific to the analyte of interest is measured. A monochromator is the line isolation device utilized for atomic absorption. All wavelengths of light enter the monochromator through an entrance slit and are subsequently divided into specific wavelengths using a prism or, more typically, a diffraction grating.

It isolates a specific emission line from other lines. It separates the resonance line from all non-absorbed lines generated by the radiation source. By adjusting the position of this dispersing device, only light of the desired wavelength enters the detector through the exit slit. Since atomic absorption is a specialized technique, very highly resolving monochromators are not required.

- VI. Detector

The monochromator’s selected light is directed onto a detector. This apparatus converts radiant energy into electrical energy, giving an electrical signal to the radiant absorbed by atoms. Typically a photomultiplier tube is used as a detector. Its function is to convert the light signal into an electrical signal proportional to the light intensity.

- VII. A-C amplifier

The electric current from the photomultiplier detector is fed to the amplifier which

amplifies the electric current several times.

- VIII. Read-out devices

It may be a simple galvanometer or chart recorder. digital display etc.

- Summary working procedure of atomic absorption spectroscopy

Light of a specific wavelength (emitted by a special type of lamp) that is capable of emitting the spectral lines corresponding to the energy required for an electronic transition from a ground state to an excited state is permitted to pass through the flame. At that time, the sample solution is drawn into the flame. Before entering the flame, the solution is dispersed into a mist of very small droplets, which evaporates in the flame to produce first the dry salt, and subsequently the salt vapor. A portion of this vapor will dissociate into atoms of the elements to be investigated.

As a result, the flame contains free unexcited atoms capable of absorbing radiation from an external source when the radiation matches exactly the energy required to transition the element from the ground electronic state to the higher excited state.

The unabsorbed radiation from the flame is then allowed to travel through a monochromator, which isolates the exciting spectral lines of the light source. The unabsorbed radiation from the monochromator is led into the detector, which is then recorded by a photodetector, the output of which is amplified and measured on a recorder. The difference in transmitted signal in the presence and absence of the test element is used to calculate absorption.