PHYSICAL PHARMACEUTICS II BP(403T) UNIT II(RHEOLOGY)

CLASS 12

Topic Name:MULTI POINT VISCOMETERS (ROTATIONAL)

The instruments are based on the fact that a solid rotating body immersed in a liquid is subjected to a retarding force due to viscous drag, which is directly proportional to the viscosity of liquid

Cup and Bob

- Couette type viscometers: Cup is rotated, the viscous drag on the bob due to sample causes to turn. The torque is proportional to viscosity of sample.

Ex. McMichael viscometer

- Searle type viscometers: Bob is rotated, the torque resulting from the viscous drag of the system under examination is measured by spring or sensor in the drive to the bob.

- Ex. Stormer viscometer

- Working: The test sample is place in space between cup and bob & allow to reach temperature equilibrium.

- A weight is place in hanger and record the time to make 100 rotations by bob, convert this data to rpm.

- This value represents the shear rate, same procedure repeated by increasing weight.

Stormer viscometer Stormer viscometer

- So then plotted the rheogram rpm Vs weights the rpm values converted to actual rate of shear and weight converted into units of shear stress, dy/cm2 by using appropriate constants.

Mathematical treatment:

- For, rotational viscometers, the relationship can be expressed as,

- Ώ=(1/η)T/4Πh(1/R 2 -1/R 2)

b	c

- Ώ=angular velocity radians/sec produced by

- T=torque in dynes-cm

- h=the depth to which bob is immersed in liquid

- Rb	= radius of the bob

- Rc=radius of the cup

- η = Kv T/ Ώ

- K=is constant for the instrument

- Ώ is a function of v the rpm generated by the weight w in grams which is proportional to T

- η = Kv w/v

- where

- v, rpm generated by weight w, in gm

- Kv is obtained by analyzing material of known viscosity in poise.

- K can be determined by analyzing an oil of known viscosity in the instrument

- Reference oils for this purpose are obtained from national bureau of standards

- The equation of plastic flow when employing stromer viscometer is

- U=Kv(W-Wf/v)

- Wf is the yield value intercept in grams

- The yield value of plastic system obtained by the use of equation f=KfXWf

- Where Kf is equal to=KX2Π/60X1/2.303log(Rf /Rb )

Plug flow

- Potential disadvantage of cup and bob viscometer

- Variable shear stress across the sample between the bob and cup

- The shear stress close to the rotating bob at relatively low rates of shear maybe sufficiently high so as to exceed yield value

- The shear stress at the inner walls of the cup could however lie below the yield value

- Materials in this zone would therefore remains as a solid plug and the measured viscosity would be in error

- Major factor determining whether or not plug flow occurs is the gap between the cup and bob

- Use of largest bob possible with a cup of definite circumference so as to reduce the gap and minimize the chances of plug flow

- In a system exhibiting plug flow in the viscometer more and more of the sample is sheared at a stress above the yield value as the speed of the bob increased

- It is only when the shear stress at the wall of the cup exceeds the yield value

,however that the system as a whole undergoes laminar, rather than plug flow and the correct plastic viscosity is obtained

- The phenomenon of plug flow is important in the flow of pastes and concentrated suspensions through an orifice

- High shear along the inner circumference of the tube aperture cause a drop in consistecy

- Facilitates extrusion of a material in the core as a plug

- However plug flow undesirable when attempting to obtain the rheogram of a plastic system with cup and bob viscometer